## Sheet1
code | name | description | unit | resolution | min | max | meaning
0 | PIP | Peak Pressure | mbar or cmH2O | 0.1 | 0 | 150
1 | MAP | Mean Pressure | mbar or cmH2O | 0.1 | 0 | 150
2 | PEEP | Positive End Expiratory Pressure | mbar or cmH2O | 0.1 | 0 | 150
3 | Ti_PSV | Inspiratory Time PSV | sec | 0.001 | 0 | 5
4 | Cdyn | Dynamic Compliance | mL/mbar or mL/cmH2O | 0.1 | 0 | 500
5 | C20_C | Overextension Index |  | 0.1 | 0 | 50
6 | R | Ventilatory Resistance | mbar/L/sec or cmH2O/L/s | 0.1 | 0 | 500
7 | MV | Minute Volume | L/min | 0.001 | 0 | 20
8 | MVresp | Share Minute Volume Respirator | % | 1 | 0 | 100
9 | VTemand | Expiratory Mandatory Tidal Volume | mL | 0.1 | 0 | 600
10 | VTemand_resp | Expiratory Mandatory Tidal Volume Respirator | mL | 0.1 | 0 | 600
11 | VTespon_pat | Expiratory Mandatory Tidal Volume Patient | mL | 0.1 | 0 | 600
12 | Leak | Leakage | % | 1 | 0 | 100
13 | RR | Respiratory Rate | 1/min | 1 | 0 | 250
14 | Trigger | Trigger Volume / Trigger flow | mL | 0.1 | 0 | 750
15 | VTimand | Inspiratory Mandatory Tidal Volume | mL | 0.1 | 0 | 600
16 | deltaP | HF Amplitude | mbar or cmH2O | 0.1 | 0 | 200
17 | VThf_emand | HF Expiratory Mandatory Tidal Volume | mL | 0.1 | 0 | 200
18 | DCO2 | Gas Transport Coefficient | mL^2/sec | 1 | 0 | 10000
19 | HFOV_freq | HOVF Rate | Hz | 1 | 0 | 20
20 | FiO2 | Franction of inspired oxygen | % | 0.1 | 21 | 100
21 | etCO2 | End-tidal CO2 | mmHg or kPa or % | 0.1
22 | Flow_demand | Demand Flow | L/min | 0.001 | 4 | 20
23 | Flow_insp | Inspiratory Flow | L/min | 0.001 | 4 | 20
24 | Flow_exp | Expiratory Flow | L/min | 0.001 | 4 | 20
25 | 25
26 | 26
27 | 27
28 | 28
29 | 29
30 | 30
31 | 31
32 | 32
33 | 33
34 | 34
35 | 35
36 | 36
37 | 37
38 | 38
39 | 39
40 | 40
41 | 41
42 | 42
43 | 43
44 | 44
45 | 45
46 | 46
47 | 47
48 | 48
49 | 49
50 | 50
51 | 51
52 | 52
53 | 53
54 | 54
55 | 55
56 | 56
57 | 57
58 | 58
59 | 59
60 | 60
61 | 61
62 | 62
63 | 63
64 | 64
65 | 65
66 | 66
67 | 67
68 | 68
69 | 69
70 | 70
71 | 71
72 | 72
73 | 73
74 | 74
75 | 75
76 | 76
77 | 77
78 | 78
79 | 79
80 | 80
81 | 81
82 | 82
83 | 83
84 | 84
85 | 85
86 | 86
87 | 87
88 | 88
89 | 89
90 | 90
91 | 91
92 | 92
93 | 93
94 | 94
95 | 95
96 | 96
97 | 97
98 | 98
99 | 99
100 | Patient_range | Patient Range |  |  |  |  | 1=Neonatal, 2=Pediatric
101 | Ventilator_mode | Ventilator mode |  |  |  |  | 0=None, 1=IPPV, 2=SIPPV, 3=SIMV, 4=SIMVPSV, 5=PSV, 6=CPAP, 7=NCPAP, 8=DUOPAP, 9=HFO, 10=O2therapy, 15=service
102 | PIP_set | Inspiratory Pressure | mbar or cmH2O | 0.1
103 | PIP_set_PSV | Inspiratory pressure PSV | mbar or cmH2O | 0.1
104 | PEEP_set | PEEP/ CPAP | mbar or cmH2O | 0.1
105 | FiO2_set | FiO2 | % | 100%
106 | Flow_insp_set | Inspiratory flow / in HFOV base flow | L/min | 0.001
107 | Slope_set | Rise Time | sec | 0.001
108 | Flow_exp_set | Expiratory Flow / Flow in O2 therapy mode | L/min | 0.001
109 | Ti_set | Inspiratory time | sec | 0.001
110 | Te_set | Expiratory Time | sec | 0.001
111 | RR_set | Rate | 1/min | 1
112 | IE_I_set | I-part TI:TE |  | 0.1
113 | IE_E_set | E-part TI:TE |  | 0.1
114 | deltaP_set | HFOV amplitude | mbar or cmH2O | 0.1
115 | Freq_set_HFOV | HFOV Rate | Hz | 1
116 | MAP_set_HFOV | HFOV MAP | mbar or cmH2O | 0.1
117 | MAP_set_HFOV_recr | HFOV lung recruitment MAP | mbar or cmH2O | 0.1
118 | Volume_limit_set | Volume Limit | mL | 0.1
119 | VG_set | Volume guarantee | mL | 0.1
120 | Not valid
121 | Term_criteria_PSV_set | Termination criteria PSV | % | 1
122 | Apnea_time_set | Apnea Time | sec | 1
123 | RR_backup_set | Backup frequency | 1/min | 1
124 | Trigger_sens_set | Trigger sensititivy |  | 1
125 | Mode_option_1
126 | Mode_option_2
127 | Powerstate | Powerstate |  |  |  |  | 0 = Network, 1=Battery
128 | Battery_rem_time | Remaining Battery time | min | 1
129 | Battery_rem_pc | Reamining Batery percent | % | 1
130 | MV_lim_high_set | MV high limit | L/min | 0.001
131 | MV_lim_low_set | MV low limit | L/min | 0.001
132 | PIP_lim_high_set | PIP high limit | mbar or cmH2O | 0.1
133 | PIP_lim_low_set | PIP low limit | mbar or cmH2O | 0.1
134 | RR_lim_set | RR limit | 1/min | 1
135 | Leakage_lim_set | Leakage limit | % | 1
136 | DCO2_lim_high_set | DCO2 high limit | mL2/sec | 1
137 | DCO2_lim_low_set | DCO2 low limit | mL2/sec | 1
138 | etCO2_lim_high_set | etCO2 high limit | mmHg or kPa or % | 0.1
139 | etCO2_lim_low_set | etCO2 low limit | mmHg or kPa or % | 0.1
140 | Measuring_unit_pressure_set | Measuring Unit pressure |  |  |  |  | 0=mbar, 1=cmH2O
141 | Measuring_unit_CO2_set | Measuring Unit CO2 |  |  |  |  | 0=mmHg, 1=kPa, 2=Vol%
142 | O2_compensation_set | O2 compensation | % | 1
143 | CO2_baropressure_set |  | mmHg | 1
144 | Device_id_cat
145 | Device_ID_var
146 | Version_HMI_maj
147 | Version_HMI_min
148 | Version_HMI_patch
149 | Version_HMI_build
150 | Version_mon_PIC_maj
151 | Version_mon_PIC_min
152 | Version_mon_PIC_patch
153 | Version_contr_PIC_maj
154 | Version_contr_PIC_min
155 | Version_contr_PIC_patch
156 | Version_aculink
157 | Flow_sensor_state | State of flow sensor |  |  |  |  | 0=off, 1=on
158 | Oxy_sensor_state | State of oxy sensor |  |  |  |  | 0=off, 1=on
159 | P_man_breath_CPAP_HFO_set |  | mbar or cmH2O | 0.1
160 | P_man_breath_duoPAP_NCPAP_set |  | mbar or cmH2O | 0.1
161 | FiO2_flush_time_set |  | sec | 1
162 | FiO2_flush_set |  | % | 1
163 | 163
164 | 164
165 | 165
166 | 166
167 | 167
168 | 168
169 | 169
170 | 170
171 | 171
172 | 172
173 | 173
174 | 174
175 | 175
176 | 176
177 | 177
178 | 178
179 | 179
180 | 180
181 | 181
182 | 182
183 | 183
184 | 184
185 | 185
186 | 186
187 | 187
188 | 188
189 | 189
190 | 190
191 | 191
192 | 192
193 | 193
194 | 194
195 | 195
196 | 196
197 | 197
198 | 198
199 | 199
200 | Alarm_susp | [5 empty] | 0=suspended, 1=active
201 | Alarm_Flat_battery | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
202 | Alarm_Checksum_ctrl_PIC | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
203 | Alarm_Checksum_monitor_PIC | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
204 | Alarm_Safety_relay_defect | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
205 | Alarm_Sens_dev_prox_pressure | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
206 | Alarm_input_pressure_blender | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
207 | Alarm_excess_pressure | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
208 | Alarm_voltage_monit | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
209 | Alarm_SPI_interface | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
210 | Alarm_DIO2_interface | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
211 | Alarm_COM_interface | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
212 | Alarm_I2C_interface | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
213 | Alarm_parallel_interface | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
214 | Alarm_serial_tem_interface | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
215 | Alarm_low_physical_memory | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
216 | Alarm_Fan_defect | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
217 | Alarm_CO2_interface | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
218 | Alarm_blender_defect | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
219 | Alarm_battery_defect | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
220 | Alarm_input_pressure_O2_supply | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
221 | Alarm_input_pressure_air_supply | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
222 | Alarm_tube_occlusion | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
223 | Not valid | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
224 | Alarm_patient_disconnected | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
225 | Alarm_ETT_blocked | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
226 | Alarm_flow_sensor_defect | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
227 | Alarm_flow_sensor_clean | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
228 | Alarm_flow_sensor_disconnected | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
229 | Oxygen_sensor_defect | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
230 | Oxygen_sensor_used_up | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
231 | Oxyen_value_divergence | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
232 | Alarm_O2_sensor_cal_error | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
233 | Alarm_MV_high | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
234 | Alarm_MV_low | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
235 | Alarm_pressure_high | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
236 | Alarm_pressure_low | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
237 | Alarm_PEEP_high | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
238 | Alarm_RR_high | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
239 | Alarm_ETT_leak_high | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
240 | Alarm_apnea | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
241 | Alarm_DCO2_high | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
242 | Alarm_DCO2_low | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
243 | Alarm_etCO2_high | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
244 | Alarm_etCO2_low | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
245 | Alarm_PIP_not_reached | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
246 | Alarm_limited_volume | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
247 | Alarm_volume_not_reached | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
248 | Alarm_power_failure | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
249 | Alarm_charge_battery_60min | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
250 | Alarm_charge_battery_30min | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
251 | Alarm_charge_battery_15min | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
252 | Alarm_nebulizer_disconnection | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
253 | Alarm_nebulizer_system_error | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
254 | Alarm_CO2_module_not_connected | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
255 | Alarm_CO2_filterline_not_connected | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
256 | Alarm_CO2_check_sampleline | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
257 | Alarm_CO2_check_airway_adapter | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
258 | Alarm_CO2_sensor_faulty | [5 empty] | 0=no alarm, 1=active alarm, 2=reversed alarm
259 | 259
260 | 260
261 | 261
262 | 262
263 | 263
264 | 264
265 | 265
266 | 266
267 | 267
268 | 268
269 | 269
270 | Ventilation_stopped | Ventilation stopped |  |  |  |  | 0=No, 1=Yes
271 | VG_state | State of volumen guarantee |  |  |  |  | 0=off, 1=on
272 | Volume_limit_state | State of volumen limit |  |  |  |  | 0=off, 1=on
273 | Ventilator_range | Ventilator range |  |  |  |  | 0=neonatal, 1=paediatric
274 | Trigger_mode | Trigger mode |  |  |  |  | 0=Volumetrigger, 1=Flowtrigger
275 | I_E_HFOV | I:E ratio HFO |  |  |  |  | 0=1:3, 1=1:2, 2=1:1
276 | Pressure_rise_control | Pressure rise control |  |  |  |  | 0=I-flow, 1=Ramp, 2=AutoIFlow
277 | HFO_recruitment | HFOV recruitment |  |  |  |  | 0=off, 1=on